Beam Size Monitors
for KEKB-ILCDR
J.W. Flanagan
KEKB-ILCDR meeting
19 Sep. 2007

---
Interferometers
Beam size currently measured by interferometer.
Resolution fundamentally limited by opening angle between slits from beam.

---
Interferometer Source Parameters
%sigma = {%lambda F} over {%pi D} sqrt{1 over 2 ln{1 over %gamma}}
D over F leslant 2 %theta_c
%theta_{c} = left ({3 c} over {%omega %rho} right )^{1 over 3}
%omega = 2 %pi c over %lambda
Note:
D = slit separation, F = distance from beam to slits.
Max slit opening angle also limited physically with current chamber to ~0.003 rad

---
Interferometers
Current interferometers cannot quite make it to the resolution needed for KEKB-ILCDR (or SuperB Low-emittance) operation.
Possible fixes (probably need a combination):
  - Increase vertical beta function at source point
  - Reduce bending radius of source magnet AND increase extraction aperture size
  - Reduce observation wavelength
    - Would gain 20% if 500 um -> 400 um.
  - Accept higher visibility: 90% -> 95% would take us from 12 um to 8 um. But error bars grow rapidly.

---
X-Ray Monitor
Used or planned to be used at ATF, CESR, Spring-8, elsewhere.
To maximize bandwidth and minimize number of components, we are considering the use of coded aperture imaging.

---
Coded Aperture Imaging
A coded aperture is a mask used to modulate incoming light.
A Fresnel zone plates is typically used as an X-ray lens
  - Requires the use of a monochromator
    - Expensive
    - Sensitive to heat load
    - Cuts available light level down drastically (1%), necessitating long exposure times
A pinhole is the simplest type of coded aperture, requiring no monochromator (good), but having a very small aperture (bad).
In 1968 R.H. Dicke (APJL, 153, L101, 1968) proposed the use of a random array of pinholes for X-ray and gamma-ray astronomy. The resulting image needs to be deconvolved back through the mask pattern to reconstruct the pattern on the sky.

---
Coded Aperture Imaging
Several improved mask designs have since been developed, most notably the Uniformly Redundant Array (URA) mask, which has the nice property that its auto-correlation is a delta function (no sidelobes), and it can achieve open aperture areas of up to 50%.
  - For a good overview and bibliography, see http://astrophysics.gsfc.nasa.gov/cai/
Several reconstruction methods are in use: inversion, cross-correlation, photon tagging (back-projection), Wiener filtering, and iterative methods such as the Maximum Entropy Method and Iterative Removal of Sources (IROS).
Coded aperture imaging is now a well-established technique in X-ray astronomy, though it has not found widespread use outside that field.
  - I have found some scattered references to uses in medical imaging, thermal neutron imaging, inertial confinement monitoring, and nuclear blast monitoring, but almost all development work seems to have been done by X-ray and gamma-ray astronomers.
I have found one reference to use of URA masks for the measurement of phase coherence of undulator radiation (J.J.A. Lin et al., “Measurement of the Spatial Coherence Function of Undulator Radiation using a Phase Mask,” PhysRevLett.90.074801), and they reference an earlier application to the same measurement of an x-ray laser (J. E. Trebes, et al., Phys. Rev. Lett. 68, 588–591 (1992).) Note: a 6 pinhole mask was tried at TRISTAN (A. Ogata et al., PAC 1989), but not with coded aperture reconstruction techniques in mind, according to Mitsuhashi.
I believe coded aperture techniques would be useful for general beam profile and position diagnostics.

---
Coded Aperture Decoding
b
f
Source
pix. size
= c(b+f)/f
Mask
min. hole size
= c
Detector
pix. size
= c(b+f)/b
Source
pix. size
= c(b+f)/f
Magnification m=(b+f)/b
Fenimore and Cannon, Appl. Optics, V17, No. 3, p. 337 (1978)

---
Modified URA Mask, Anti-mask, and Cross-correlation
Image is encoded using mask and decoded using anti-mask, where cross-correlation between mask and anti-mask is delta function.
Pixel transparency determined by Jacobi function:
  - Is (pixel index)%DIM == (i*i)%DIM for any 1<i<DIM?
    - Yes/No->Open/Closed.
    - 2-D case based on inverse XOR of both indices.
Note: Fresnel zone plates can in principle also be used as coded apertures. (Barrett, H.H., Horrigan, F.A.: 1973, Appl. Opt., 12, 2686)

---
Examples
As an illustrative example, here is a simulation of a 13x13 pixel source image, projected through a 13x13 Modified URA mask onto a 26x26 CCD.
  - The image represents a beam with σy=5 μm (typical of ILC Damping Ring study mode, or SuperB super-low emittance mode) and σx=10μm, with minimum mask pinholes 4 μm on a side.
  - With a 5:1 magnification factor (e.g., mask 6 meters downstream of source, and CCD 24 meters downstream of mask), the CCD pixels would be 25 μm on a side, which is about the size of the x-ray CCD in use at the ATF. The source resolution elements would be 5 μm on a side.
The reconstruction method used is direct decoding.
In the second case, a random scattering of 10% noise has been added to the CCD image, which has then been reconstructed via decoding.

---
URA Mask, decoding via anti-mask
Source Image
Mask
CCD Image
Reconstructed Image and Profiles

---
URA Mask, Decoded, with 10% noise on CCD
Source Image
Mask
CCD Image
Reconstructed Image and Profiles

---
LER X線モニタービームライン（１）

---
X-Ray Source Bend (B2P.53)

---
LER X線モニタービームライン（２）
Weak bend
光モニター光源
ヨークを削る
アンテチェンバー
に変更（w/ photon stop)
detector
X-ray
1m
大口径 Q に変更
Normal bend
X線モニター光源
30m

---
X-Ray Source & Beamline Parameters

---
X-ray attenuation lengths

---
Energy dependence of attenuation and scattering
Compton, Rayleigh scattering start to become significant above ~ 20 keV

---
Assumptions
Beam -> mask: 6 m
Mask -> CCD: 24 m => 5x magnification
Be window thickness: 1 mm
Al filter/window thickness: 0.5 mm
Mask: 4 um-thick Tantalum on 2 um-thick SiC
  - Outer size: 0.04 mrad (V) x 5*0.04 mrad (H) (0.24 x 1.2 um @ 6 m)
    - Useful vertical size limited by critical angle
CCD quantum efficiency: 10%
  - ~true for direct detection CCD
  - higher for fluorescent screen

---
Basic layout
LER Beam
X-ray CCD
Coded
Aperture
Mask
Be
Window
Photon
Stop
Gate
Valve
Aluminum
window
6 m
24 m

---
KEKB-ILCDR mode
F = 1.33x10^{13} E(GeV)^2 I(A) H_2(%omega/%omega_c)~;~ H_2(y) = y^2 K_{2/3}^2(y/2)
Flux through mask holes
K.J. Kim, AIP Conf. Proc. 184 (1989)
Flux through mask
shadow region (Ta)
%theta_c = 0.58 over %gamma [ %omega over %omega_c ]^{-1/2} ~;~%omega >> %omega_c

---
KEKB-ILCDR mode
Gamma = 4.500978e+03
Critical energy = 1.945305e+00 keV
Total source power = 2.812707e-02 kW/mrad
Flux from source: 1.62907e+17 photons/s/mr^2/Ampere
Flux after 0.1 cm Be: 1.91572e+16 photons/s/mr^2/Ampere
Flux after 0.05 cm Al: 2.12417e+14 photons/s/mr^2/Ampere
Flux after 0.0004 cm Ta: 7.37955e+13 photons/s/mr^2/Ampere
Flux after 0.0002 cm SiC: 2.10648e+14 photons/s/mr^2/Ampere
Flux after 10 cm Air: 2.03274e+14 photons/s/mr^2/Ampere
Flux through 0.008 mr^2 mask: 1.6262e+12 photons/s/Ampere
Flux/turn 1.6262e+07 photons/turn/Ampere
Flux/mA/bunch 16262 photons/turn/mA/bunch
Detected signal 1626.2 photons/turn/mA/bunch
Detected background 566.748 photons/turn/mA/bunch
On-axis power from source: 0.0736069 kW/mr^2/Ampere
On-axis power after 0.1 cm Be: 0.0190535 kW/mr^2/Ampere
On-axis power after 0.05 cm Al: 0.000446705 kW/mr^2/Ampere
On-axis power after 0.0004 cm Ta: 0.000164321 kW/mr^2/Ampere
On-axis power after 0.0002 cm SiC: 0.000443388 kW/mr^2/Ampere
On-axis power after 10 cm Air: 0.000429463 kW/mr^2/Ampere

---
But, diffraction effect is not small, as Mitsuhashi points out
x-ray: 5 keV
URA mask: 23x23
Hole size: 2.4 um
Distance from mask to camera: 3 m
Diffraction calculated using Zemax
This can in principle still be reconstructed if we know the spectrum, using iterative methods such as maximum entropy.

---
Vertical-only mask: 1x31
Much faster reconstruction when using iterative methods
(1-D vs 2-D problem)
1-D URA Mask
Autocorrelation

---
Vertical-only mask: 1x31, 4 um min. aperture
Irradiance as function of photon energy. Mask->detector = 24 m
6.2 keV
12.4 keV
24.8 keV
Averaged over spectrum

---
URA 1x31 x 4 um; Decoding; Beam sigy=5 um, spec. 5-30 keV; No Noise
Source Image
CCD Image
Mask Irradiance w/diffraction
Reconstructed Image and Profile

---
URA 1x31 x 4 um; Max. Ent. reconstruction; Beam sigy=5 um, 5-30 keV; No Noise
Source Image
CCD Image
Mask Irradiance w/diffraction
Reconstructed Image and Profile

---
URA 1x31 x 4 um; Max. Ent. reconstruct.; Beam sigy=5 um, 5-30 keV; 10% Noise
Source Image
CCD Image
Mask Irradiance w/diffraction
Reconstructed Image and Profile

---
Other hardware
In discussion with G. Varner of UH (Belle pixel detector person) on high-speed detector/readout system.
  - 6 GS/s -> slice bunch
Costs (very rough!):
  - Extraction chamber: 4000 man-yen ($400,000 at 100 yen/$)
  - Mask: 250 man-yen * N (NTT Adv. Tech.)
  - Detector: 750 man-yen?
  - Hole in magnet for extraction line: 250 man-yen
  - Extraction line, Be window, Al window, gate valve, photon stop: 1750 man-yen?
  - Total: ~7500 man-yen?

---
Conclusion
Coded Aperture Imaging seems to be a realistic possibility for x-ray beam profile and position monitoring.
Work needs to be done on selecting an appropriate mask pattern and reconstruction method.
  - URA decoding is very fast, but cannot handle diffraction effects, for which we need iterative methods.
But in principle, a very simple and relatively low-cost system might be able to be constructed from, say, a beryllium window in the beam pipe at a source bend, a mask, and an x-ray CCD, plus perhaps an aluminum filter if needed to reduce power.
  - Mitsuhashi offers use of x-ray tube at PF to start testing
  - Design with view towards KEKB ILC mode or SuperB low-emittance mode